Regulation on Metrological Control

of  Non-automatic Weighing Instruments.

CHAPTER I

Scope and definitions.

Article 1.

Scope.

The scope of the present Regulations is non-automatic weighing instruments which are subject to the scope of Regulations no 616/2002 on non-automatic weighing instruments, hereafter named weighing instruments.

Weighing instruments which are subject to verification according to Article 5 shall comply with further requirements to receive regular verification according to the present regulations.

The verification of non-automatic weighing instruments is divided into two areas of verification. The verification of non-automatic weighing instruments is more specifically:

a. Up to and including 3000 kg.

b. Above 3000 kg.

Article 2.

Definitions.

The following definitions shall be used for weighing instruments  according to para.1 of Article 1.

- Internal inspection in the present regulation is used in the meaning included in para. 2 of Art. 14 of Act no 91/2006 on Measurements, Measurement Standards and Official Weighers and according to further provisions in the present Regulations.

- Seals are usually buttons or labels according to Regulation no 955/2006 on verification symbols and markings of measuring instruments subject to verification but may also be software seals where a password must be entered to gain access to adjustments and all such access is recorded. Sealing shall be in accordance with instructions from the manufacturer.

- Verification is a procedure to ensure and confirm formally that a weighing instrument complies with all requirements made in Acts and Regulations. Verification is carried out by inspection, labelling and/or issuing a certificate and is usually completed by the sealing of access to adjustments.

- Maximum permissible error is the highest numerical value for error permissible according to the present Regulations by verifications of non-automatic weighing instruments.

- A weighing instruments supplier is a producer, authorised representative, seller or distributor of weighing instruments subject to para. 1 of Art. 1, cf. Art. 8 of Act no 91/2006 on Measurements, Measurement Standards and Official Weighers.

- A non-automatic weighing instrument is a weighing instrument which requires the intervention of the operator during weighing.

- Conformity assessment is when the producer certifies that the weighing instruments are in accordance with the essential requirements of Annex I in Regulations no 616/2000 with procedures which consist of one or more of the modules EC type examination, EC declaration of type conformity, EC verification and EC unit verification. The conformity assessment is completed with  CE marking and a declaration of conformity in writing from the manufacturer.

- A notified body is an independent testing-, certification-  or inspection body notified  by the authorities to undertake conformity assessments in accordance with Regulations no 616/2000 on non-automatic weighing instruments.

- A weighing instrument is a measuring instrument serving to determine the mass of a body by using the action of gravity on  that body. A weighing instrument may also serve to determine other mass-related magnitudes such as counts, quantities, parameters or characteristics.

In addition to the definitions above the technical terms and symbols shall be explained with reference to their use in Regulations no 616/2000 on non-automatic weighing instruments.

CHAPTER II

Older weighing instruments, first use and labelling, verification duty.

Article 3.

Continued use of older types of weighing instruments

Weighing instruments which do not meet the requirements of Regulations no 616/2000 cf. para. 1 of Art. 1, may be in continued use while they weigh correctly if they have been properly placed on the market and put into use before January 1. 2004

Older types of weighing instruments shall pass the same requirements for precision as weighing instruments according to Regulations no 616/2000 on non-automatic weighing instruments.

Article 4.

First use and labelling.

Weighing instruments fulfilling the requirements of Regulations no 616/2000 shall have conformity assessment, correct markings and seals and they shall then be considered to have verification for first use. They shall receive a special marking for new weighing instruments according to Regulations no 955/2006 onVerification Symbols and Markings of Weighing Instruments Subject to Inspection.

When para. 5 of Annex II in Regulations no 616/2000 applies in regard to installation or gravity and a declaration of conformity is completed at the place of use of the weighing instrument, the first use is not permitted until the declaration of conformity is fully completed and a declaration of conformity assessment in writing from the manufacturer is available.

Article 5.

Weighing instruments subject to verification.

Verification shall be obtained for weighing instruments used for the following purposes:

a.	Determination of mass for business transactions.

b.	Determination of mass for the calculation of a toll, tariff, tax, bonus, penalty, remuneration, indemnity or similar type of payment.

c.	Determination of price on the basis of mass for the purposes of direct sales to the public and the making-up of prepackages

d.	Weighing according to Regulation no 224/2006 on Weighing and Recording of Fish Catches in which weighing shall be carried out on a verified weighing instrument even though it may not be covered by section a.

If doubt arises on the obligation to obtain verification the matter is decided by the Consumer Agency and it can also award exceptions in cases when use of a weighing instrument constitutes a low risk of harming any interests. Other authorities may also require an obligation to obtain verification e.g. for the weighing of fish catches.

CHAPTER III

Responsibility of parties.

Article 6.

Requirements of suppliers.

A supplier of weighing instruments shall not market, sell or deliver weighing instruments according to para. 1 of Art. 1 unless their conformity assessment is completed, they have received a written declaration of conformity of the manufacturer, are sealed and properly marked according to Regulations no 616/2000 on non-automatic weighing instruments, but cf. however. para. 2 of Art.4.

A supplier of weighing instruments shall inform the buyer about the accuracy class, maximum and other technical parameters of the weighing instruments and whether they are likely to be suitable for the intended use.

Article 7.

Requirements of owners of weighing instruments.

The owner may not set up or have set up a weighing instrument for use according to para.1 of Art. 1 unless the instrument is properly marked and sealed.

The owner of a weighing instrument which is subject to verification is responsible for receiving regular verification from an  authorized testing body according to Regulations 956/2006 on the Mode of Work of Those Undertaking Reverifications commissioned by the Consumer Agency.

If an owner learns of errors in use of a weighing instrument amounting to double the error limits that are used for reference for new weighing instruments or if for any other reason the verification has become invalid, according to para. 3 of Art. 16, he shall be responsible for taking the weighing instrument out of use until it has been reverified.

The owner of a weighing instrument shall adhere to the conditions of the manufacturer for conditions of use and permitted operating conditions. The requirements in Regulations no 616/2000 on non-automatic weighing instruments and in the present Regulations shall be kept in mind.

Article 8.

Requirements of service agents.

A weighing instrument which is subject to verification according to Art. 5 but does not have a valid verification according to Art. 16 shall be reverified prior to being put into use again, but cf. the permit of the servicing agent in Art. 16.

A service agent who has carried out any operation which requires reverification according to para.1 shall immediately notify the owner of this.

CHAPTER IV

Requirements of weighing instruments in use.

Article 9.

Market surveillance.

Weighing instruments are subject to market surveillance according to Regulation no 616/2000 independently of their verification obligations.

Article 10.

Maximum permissible errors.

Maximum permissible errors of weighing instruments are according to Table 1 which is identical to Table 3 in Regulations no 616/2000 on non-automatic weighing instruments, which is in force for new weighing instruments:

TABLE 1

Maximum permissible errors at verification.

| Load | Maximum permissible error |
| --- | --- |
| Class I | Class II | Class III | Class IIII |  |
| 0  m  50000 e | 0  m  5000 e | 0  m  500 e | 0  m  50 e | ±0,5 e |
| 50000 e < m  200000 e | 5000 e < m  20000 e | 500 e < m  2000 e | 50 e < m  200 e | ±1.0 e |
| 200000 e < m | 20000 e < m  100000 e | 2000 e < m  10000 e | 200 e < m  1000 e | ±1.5 e |

Article 11.

Other requirements.

In testing for verification, weighing instruments shall fulfil the requirements of Regulation no 616/2000 on non-automatic weighing instruments, however, cf. Art. 3 on older weighing instruments and Art. 14 on minimum markings.

The support of weighing instruments shall be suitable for each type of instrument and its use, it shall be stable, level and free of dirt that might affect the measurements of the weighing instrument.

The ambient temperature shall be suitable for the weighing instrument and it is not permissible to verify a weighing instrument if the ambient temperature is outside limits for which the instrument is designed.

If a weighing instrument is used for direct sales to the public it shall also comply with requirements according to para. 14 in Annex I of Regulation no 616/2000.

A weighing instrument shall not be verified or used unless disturbance due to weather conditions or the climactic conditions have a minimal effect on the results, so that it may be concluded that the requirements for maximum permissible errors are fulfilled.

Article 12.

Weighing instruments to be taken out of use.

If it becomes clear that weighing instruments do not fulfil the requirements of this Chapter they shall be taken out of use and shall not receive reverification.

In spite of the requirements of this Chapter it is permitted to use a weighing instrument until the next verification in case its error does not exceed the double maximum permissible error.

CHAPTER V.

Rules on verification of weighing instruments.

Article 13.

Weights used for verification.

Weights used for verification of weighing instruments shall fulfil the requirements of existing regulations on weights and be suitable to the conditions and the weighing instrument being validated.

Weights in at least accuracy class MI for verification of weighing instruments in class III, and class FI for verification of weighing instruments in class II shall be used. The accuracy classes MI and FI of weighs are defined in Regulations no 136/1994.

The requirement for total uncertainty for verification in regulation no 956/2006 on the procedures for parties responsible for verification on behalf of the Consumer Agency may lead to the conclusion that the level of precision as stipulated in para.2 is not satisfactory.

Article 14.

Testing requirements for verification.

Inspections or tests shall be made to ascertain if a weighing instrument complies with requirements according to Chapter IV. Tests of weighing instruments shall be carried out in conditions similar to those of normal use, as far as possible.

An inspection shall be of markings, seals and information on the weighing instrument and its display, the conditions cf Art. 11,. and if its use is suitable to the weighing instrument.

Metrological characteristics shall be tested and in order that a measuring instrument passes testing the errors shall not exceed the maximum permissible errors according to Table 1. This shall apply to each feature that is being tested.

The total measuring range of the weighing instrument shall be tested, especially the minimum and maximum and the points where the maximum permissible error changes and perform repeatability and eccentricity tests. If indications point to disturbances in weighing, the influence of possible error sources shall be examined, such as draft, a strong electrical current, a magnetic field or absorbed radiation from heat, light or mobile telephones.

The verification of older weighing instruments may be permitted, cf Art. 3., even though markings are not in accordance with the requirement of these Regulations, on the condition that minimum markings appear. Minimum markings consist of:

1.	Type of weighing instrument.

2.	Maximum capacity

3.	Minimum capacity.

4.	Verification scale interval.

5.	Serial number.

Article 15.

Verification.

A weighing instrument may be verified if it has passed inspection and testing according to Art. 14. Verification is confirmed by attaching a verification label to the weighing instrument and a report is made on its testing and verification. If the weighing instrument has devices which the owner is not permitted to remove or adjust this shall be ensured by affixing a seal prior to attaching the verification label to the weighing instrument.

The testing body responsible for verification shall not carry out repairs on a measuring instrument. However, when the errors of a weighing instrument exceed the permitted limits the verification body is permitted to make adjustments to the instrument if such adjustments are made in accordance with the manufacturer's instructions or approved methods are used when these instructions are not available.

Article 16.

Verification periods and approval of the service provider.

The verification period of weighing instruments is two years.

The verification period for new weighing instruments is based on the year on the CE marking. Thus, a weighing instrument shall be verified before the end of the second year following the year on the CE marking.

Verification becomes invalid even though the period of verification has not ended if:

- The weighing instrument breaks down,

- the seal is broken,

- repairs are made on the weighing instrument which can affect the results of its measurements,

- a weighing instrument is moved to another gravity zone than which it was verified or adjusted for,

- the error exceeds the double maximum permissible error according to Table1 in Art. 10.

Despite the provision of para. 3, the service provider approved by the Consumer Agency may be permitted to break the seal and carry out minor repairs without voiding the verification. The service provider shall seal the weighing instrument again with his identification and send a report to the Consumer Agency about what has been done.

The approval of the service provider is subject to conditions and decisions of the Consumer Agency following an examination of his qualifications and quality of work and to ensure that results of measurements will remain within the maximum permissible errors in use.

The Consumer Agency issues a written permit to a service provider with further provisions and limitations on permitted repairs. If a service provider does not meet the conditions which are the requirements for the approval of the Consumer Agency, it may revoke and invalidate the permit of a service provider in accordance with this Article.

Article 17.

Reports.

Reports to the Consumer Agency shall be in accordance with Art. 11 of Regulations no 956/2006 on the procedures for parties responsible for verification on behalf of the Consumer Agency.

The following information shall appear in test reports put forward as a basis for verifications and these shall be accessible to the staff of the Consumer Agency:

1.	Error at zero indication.

2.	Weighing characteristics and errors in the areas tested.

3.	Eccentricity error.

4.	Findings of repeatability test.

5.	Tare error.

6.	Environment and conditions.

CHAPTER VI.

Internal inspection.

Article 18.

General conditions concerning internal inspection.

Internal inspection is permissible in place of verifications upon approval of the Consumer Agency according to the provisions of the present Regulations cf. Art. 14 of Act no 91/2006 on Measurements, Measurement Standards and Official Weighers.

Internal inspection shall ensure that all weighing instruments subject to internal inspection fulfil the same requirements regarding precision as weighing instruments which are verified regularly.

Testing of weighing instruments under inspection shall be carried out according to Chapter V.

When internal inspection is used in place of verification, a record shall be kept of all information providing an overview of all measuring instruments subject to the coverage of the present regulations. The Consumer Agency shall be supplied with relevant information on the weighing instruments from this record.

The internal inspection system shall be recorded in full in the inspection or quality management handbook of the owner.

Article 19.

Application and procedure for internal inspection.

The owners of weighing instruments may request the approval of the Consumer Agency for a certified quality management system, extending to internal inspection of weighing instruments, to be accepted in place of verification.

The application shall be accompanied by all material from the quality management handbook which is relevant to internal inspection of non-automatic weighing instruments.

In the application process for approval of internal inspection the applicant shall supply all information on the arrangement of internal inspection, data on the management of the system and necessary information that must be evaluated and give indications for improvements on the projected internal inspection system.

The Consumer Agency makes a decision on the approval on the basis of submitted documents and the improvements which the applicant has carried out, if applicable. If an applicant for internal inspection fulfils all the conditions determined by the Consumer Agency it issues a permit for the use of internal inspection in place of verifications on acceptance of payment of a licencing fee according to a tariff determined by the Minister.

If the Consumer Agency turns down an application this shall be supported by reasoning in writing.

If the internal inspection does not fulfil the conditions any longer, the Consumer Agency shall revoke the permit for internal inspection of weighing instruments.

Article 20.

Annual inspection and reporting.

The owner of weighing instruments who has received the approval of he Consumer Agency for internal inspection cf Art. 19 and 20, shall submit an annual report to the Consumer Agency and no later than March 1st for the preceding year, documenting all examinations and inspections, results of inspections and decisions that have been taken concerning weighing instruments.

The report shall include a statistical summary of all types of breakdowns, faults, errors, repairs, transport and broken seals of weighing instruments. Furthermore, a survey of test results shall be given cf. Art. 17. The classification of deviations and the contents of the report shall be determined by the Consumer Agency.

The Consumer Agency shall confirm an extension of the duration of a permit for internal inspection, no later than on June 1 upon receiving an annual inspecion fee according to a tariff determined by the Minister.

CHAPTER VII.

Supervision, market surveillance, administrative resorts et al.

Article 21.

The task of the Consumer Agency.

The Consumer Agency supervises the abidance by the provisions of the present Regulations and undertakes market surveillance in conformity with Directive 2004/22/EC on Measuring Instruments and the provisions of Act no 91/2006 on Measurements, Measurement standards and Official Weighers.

Article 22.

Administrative Resorts.

The implementation of inspection of measuring instruments, judicial solutions and penalties for violations of the provisions of the present Regulations are subject to the provisions of Act no 91/2006 on Measurements, Measurement standards and Official Weighers.

Article 23.

Fees on Account of Internal Inspection.

The Consumer Agency collects a licensing fee and an annual inspection fee on account of internal inspection, cf. Art 19 and 20 and as per a tariff laid down in conformity with the provisions of Art. 32 of Act no 91/2006 on Measurements, Measurement Standards and Official Weighers.

CHAPTER VIII

Appeal and means of complaints.

Article 24.

Appeal and Means of Complaints

Decisions made by the Consumer Agency on the basis of the present Regulations can be appealed to the Consumer Appeals Committee which acts on the basis of Art. 4 of Act  62/2005 on the Consumer Agency and Consumer Spokesman.

The decision of the Consumer Agency can not be submitted to Courts of Law until a ruling from the Consumer Appeals Committee is available

CHAPTER IX

Penalties.

Article 25.

Penalties.

A party who deliberately or inadvertently violates the provisions of the present Regulations shall be subject to monetary fines or imprisonment for up to six months in case of gross violations

A party giving incorrect or misleading information in connection with decisions or who does not reveal issues of importance for disclosing a Case under the present Regulations shall be subject to fines, unless a heavier penalty applies under Laws.

Those violating the provisions of Chapters III  and VI shall be subject to fines.

A fine under the present Regulations may be imposed on legal persons in conformity with provisions of Penal Laws on fines for legal persons.

CHAPTER X

Entry into force, alignment et al.

Article 26.

Entry into Force and Alignment.

The present Regulations are laid down with support in Act no 91/2006 on Measurements, Measurement Standards and Official Weighers.

At the same time, the Regulations no 793/2002 on the verification of non-automatic weighing instruments and no 958/2006 on amendment of the Regulations the verificatin of non-automatic weighing instruments and no 34/2008 on amendment of the Regulations on Regulations the verificatin of non-automatic weighing instruments are repealed. .

Ministry of Business Affairs	.... ... 2008

Björgvin G. Sigurðsson

Áslaug Árnadóttir